## Alarm list
No. | AlarmID. | Output Alarm Condition | Subsystem | Input Alarm Descriptor | Priority | Classification | Group | Alarm Reaction CURRENT OUTPUT | Alarm Reaction DIGITAL OUTPUT | Mask State | HMI Description | HMI Recovery Action | Remap | Cause(For service) | Remedy(For service) | Ref. from
1 | 0 | VORTEX_SENSOR_FAILURE | Coordinator | CDT_VORTEX_SENSOR_CIRCUIT_FAILURE | 213 | FAILURE | HW_STATUS_SENSOR | general alarm | Collective Alarm | no | Sig. Sensor Fault | Signal sensor fault. Contact ABB Service | PV_SENSOR_FAILURE | Vortex signal path circuit error | Replace piezoelectronic sensor or contact ABB Service. | Merrick
2 | 0 |  | Coordinator | CDT_VORTEX_SENSOR_OPEN | [9 empty] | Vortex signal sensor open error |  | Merrick
3 | 0 |  | Coordinator | CDT_VORTEX_SENSOR_SHORT | [9 empty] | Vortex signal sensor short error |  | Merrick
4 | 0 |  | Coordinator | CDT_VORTEX_SENSOR_DRIFT | [9 empty] | Vortex signal sensor drift error |  | Merrick
5 | 1 | TEMP_SENSOR_FAILURE | Coordinator | CDT_TEMP_SENSOR_UPPER_LIMIT | 212 | FAILURE | HW_STATUS_SENSOR | general alarm | Programmed Alarm | Single mask | Int. T Sesnor Fault | Internal temperature sensor fault. Contact ABB Service | SV_SENSOR_FAILURE | Temperature sensor over upper limit | Replace Sensor board or contact ABB Service. | Merrick
6 | 1 |  | Coordinator | CDT_TEMP_SENSOR_LOWER_LIMIT | [9 empty] | Temperature sensor below lower limit |  | Merrick
7 | 2 | VIBRAT_SENSOR_FAILURE | Coordinator | CDT_VIBRAT_SENSOR_CIRCUIT_FAILURE | 211 | FAILURE | HW_STATUS_SENSOR | general alarm | Collective Alarm | no | Vbr. Sesnor Fault | Vibration sensor fault. Contact ABB Service | TV_SENSOR_FAILURE | Vibration signal circuit error | Replace piezoelectronic sensor or contact ABB Service. | Merrick
8 | 2 |  | Coordinator | CDT_VIBRAT_SENSOR_OPEN | [9 empty] | Vibration signal sensor open error |  | Merrick
9 | 2 |  | Coordinator | CDT_VIBRAT_SENSOR_SHORT | [9 empty] | Vibration signal sensor short error |  | Merrick
10 | 2 |  | Coordinator | CDT_VIBRAT_SENSOR_DRIFT | [9 empty] | Vibration signal sensor drift error,Sensor degradation |  | Merrick
11 | 3 | ANALOGINPUT_OUT_OF_RANGE | Analog input | AI_ALM_OUT_OF_RANGE | 106 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Group mask | AI Out of Range | Check AI Current | FV_SENSOR_FAILURE
12 | 4 | FLOWRATE_REACH_QMAX | MVM | MVM_ALM_FLOWRATE_REACH_QMAX | 114 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | Programmed Alarm | Single mask | Max Alarm Flowrate | Flowrate greater than the max. Check Limits Setting | PV_SENSOR_OUT_LIMS | Flowrate reach Max Alarm Limits. | check limits setting | SRS
13 | 4 |  | SVM | SVM_ALM_FLOWRATE_REACH_QMAX
14 | 5 | PROCESS_TEMP_REACH_TMAX | Coordinator | CDT_ALRAM_T_MAX | 112 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Single mask | Max Alarm Int. Temp | Process temp. greater than the max. Check Limits Setting | SV_SENSOR_OUT_LIMS | Process Temperature reach Max Alarm Limits. | check limits setting | SRS
15 | 6 | ANALOGINPUT_CUT_OFF | Analog input | AI_ALM_CUT_OFF | 107 | OFF_SPECIFICATION | OPERATING_CONDITION | 4mA | No response | Group mask | AI Out of Range | Check AI Current | TV_SENSOR_OUT_LIMS
16 | 7 | PRESSURE_REACH_PMAX | Mapper | MAPPER_ALM_PRESSURE_REACH_MAX | 102 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Group mask | Max Alarm Pressure | Pressure greater than the max. Check Limits Setting | FV_SENSOR_OUT_LIMS | Pressure reach Max Alarm Limits. | check limits setting | Carsten
17 | 10 | FLOWRATE_REACH_QMIN | MVM | MVM_ALM_FLOWRATE_REACH_QMIN | 113 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | Programmed Alarm | Single mask | Min Alarm Flowrate | Flowrate lower than the min. Check Limits Setting | PV_SENSOR_OUT_RANGE | Flowrate reach Min Alarm Limits. | check limits setting | SRS
18 | 10 |  | SVM | SVM_ALM_FLOWRATE_REACH_QMIN
19 | 11 | PROCESS_TEMP_REACH_TMIN | Coordinator | CDT_ALRAM_T_MIN | 111 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Single mask | Min Alarm Int. Temp | Process temp. lower than the min. Check Limits Setting | SV_SENSOR_OUT_RANGE | Flowrate reach Min Alarm Limits. | check limits setting | Carsten
20 | 12 | CURRENT_OUT_SATURATED | Current Out | CURR_DIAG_SPANLIMIT_LO | 108 | OFF_SPECIFICATION | OPERATING_CONDITION | Saturated value | No response | group mask | Current out saturated | Check current out saturation limits. | TV_SENSOR_OUT_RANGE
21 | 12 |  | Current Out | CURR_DIAG_SPANLIMIT_HI |  |  |  | Saturate Hi value
22 | 13 | PRESSURE_REACH_PMIN | Mapper | MAPPER_ALM_PRESSURE_REACH_MIN | 101 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Group mask | Min Alarm Pressure | Pressure lower than the min. Check Limits Setting | FV_SENSOR_OUT_RANGE | Pressure reach Min Alarm Limits. | check limits setting | Carsten
23 | 16 | SNR_FAILURE | Coordinator | CDT_SNR_FAILURE | 210 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Bad SNR | Increase the Flow. Check Limits Setting | FE_BOARD_FAILURE | SNR too low | Replace piezoelectronic sensor or contact ABB Service. | Merrick
24 | 17 | FE_NV_FAILURE | Coordinator | CDT_FE_NV_FAILURE | 209 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Sensor NV Error | Replace the Front end Board |  | CRC Check and access failure Frontend board EEPROM faults | Replace frontend board or contact ABB Service. | Merrick,FEP300
25 | 18 | SENSOR_CAL_STATUS | Coordinator | CDT_ALARM_CB_UNCALIBRATED | 51 | MAINTENANCE | OPERATING_CONDITION | Current value | No response | Group mask | Sensor Not Calibrated | Sensor not calibrated. Contact ABB Service | FE_NV_WARNING | Sensorand/or FEB not calibrated |  | FEP-300
26 | 19 | FE_SYNC_MISSING | Coordinator | CDT_SYNC_MISSING | 214 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Sync. Signal Error | Contact ABB Service | FE_BOARD_NOT_DETECTED | Synchronous fail Frontend board circuit faults | Replace frontend board or contact ABB Service. | Merrick,FEP300
27 | 20 | FE_BOARD_COMM_ERROR | FrontendInterface | FE_BOARD_COMM_ERROR | 215 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Sensor Comm Error | Sensor and transmitter comms error. Contact ABB Service |  | No Handshake Communication, CRC error, Time out,Wire open or short. | Check the connection between front end and communication board or contact ABB Service. | Merrick
28 | 23 | CB_NV_FAILURE | Coordinator | CDT_ALARM_CB_NV_ERROR | 207 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Transmitter NV Error | Replace the Communication Board |  | CRC error and Access failure |  | 4-wire common
29 | 24 | AI_SPI_COM_FAIL | Analog input | AI_ALM_SPICOM_FAIL | 202 | CHECK_FUNCTION | CONFIG_STATUS | general alarm | Collective Alarm | no | AI Comm Error | Analog input comms error. Check IO Board | CB_NV_WARNING | Voltage input out of range or fault of AD chip | Check voltage input or change transmitter
30 | 25 | DIGITAL_OUT_CUTOFF | Digital Out | DIGITALOUTPUT_PULSENUMMAXALARM | 103 | OFF_SPECIFICATION | OPERATING_CONDITION | Current Value | No response | Group mask | Pulse Output Cutoff | Pulse output cut off. Check Pulse Output Config | NV_STORAGE_ACTIVE |  | Calculate max pulse factor and notify the end user during configuration if pulse frequecy exceeds the the maximum. | DiagnosisInfo from Kevin, FEP300
31 | 26 | RE_OUT_LINEAR_RANGE | SVM | SVM_ALM_RE_OUT_LINEAR_RANGE | 109 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Group mask | Re. out of Range | Re. out of linear range. Check Limits Setting | NV_CONCISTENCY_CHECK | RE out of linear range | check limits setting | SRS
32 | 30 | STEAM_STATUS_MISMATCH | MVM | MVM_ALM_STEAM_STATUS_MISMATCH | 116 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Group mask |  |  |  | The steam type is no longer as configured in “Device Setup->Plant/Customized->Compensation Setting->Steam Type” | Check Steam Type Setting in “Device Setup->Plant/Customized->Compensation Setting->Steam Type”
33 | 31 | MAINTENANCE_MATURE | Coordinator | CDT_ALARM_CB_MAINTENANCE | 52 | MAINTENANCE | OPERATING_CONDITION | Current Value | No response | Group mask | Maintenance Warning | Contact ABB Service |  | ADC Voltage out of range | Replace frontend board or contact ABB Service. | Merrick,FEP300
34 | 31 |  | Coordinator | CDT_FE_MAINTENANCE_MATURE
35 | 32 | CB_INSUF_INPUT_VOLTAGE | Electronic Services | ELETTRONICSERVICES_ALARM_POWER_SUPPLY_TOO_LOW | 53 | MAINTENANCE | OPERATING_CONDITION | Current Value | No response | Group mask | Voltage warning | Internal supply voltage warnning. Contact ABB Service |  | Internal voltage supply for the transmitter is defective (+/- 5, 3.3 or 24 V). | Replace the electronic unit or define the threshold. | SRS,FEP300
36 | 32 |  | Electronic Services | ELETTRONICSERVICES_ALARM_POWER_SUPPLY_TOO_HIGH | [11 empty] | SRS,FEP300
37 | 33 | AMBIENT_TEMP_REACH_TMIN | Electronic Services | ELETTRONICSERVICES_ALARM_AMBIENT_TEMP_TOO_LOW | 104 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Single mask | Min Alarm Housing Temp | Housing temp lower than the min. Check Limits Setting |  | Ambient Temperature Lower than min Alarm Limits. | check limits setting | SRS
38 | 33 |  | Current Out | CURR_DIAG_TCOMP_SATURATED_LO |  | OFF_SPECIFICATION
39 | 34 | AMBIENT_TEMP_REACH_TMAX | Electronic Services | ELETTRONICSERVICES_ALARM_AMBIENT_TEMP_TOO_HIGH | 105 | OFF_SPECIFICATION | OPERATING_CONDITION | Current value | No response | Single mask | Max Alarm Housing Temp | Housing temp greater than the max. Check Limits Setting |  | Ambient Temperature reach Max Alarm Limits. | check limits setting | SRS
40 | 34 |  | Current Out | CURR_DIAG_TCOMP_SATURATED_HI |  | OFF_SPECIFICATION
41 | 35 | LOW_FLOWRATE_CUTOFF | SVM | SVM_ALM_LOW_FLOWRATE_CUTOFF | 110 | OFF_SPECIFICATION | OPERATING_CONDITION | 4mA | Programmed Alarm | Single mask | Flowrate Cutoff | Check Low Flow Cutoff Setting |  | Flowrate is Lower than Cut off Flow. | check cut of value | 4-wire common
42 | 36 | FLOWRATE_BG_103 | MVM | MVM_ALM_FLOWRATE_GT_103 | 115 | OFF_SPECIFICATION | OPERATING_CONDITION | programmed alarm | Collective Alarm | Single mask | Flowrate > 103% | Flowrate > 103%. Check Flowrate. Check Limits Setting |  | Flowrate is bigger than 103% | check limits setting | FEP300
43 | 36 |  | SVM | SVM_ALM_FLOWRATE_GT_103
44 | 37 | DATA_SIMU_ALARM | Coordinator | CDT_ALARM_CB_SIMUDATA | 151 | CHECK_FUNCTION | CONFIG_STATUS | Current Value or simulated value (Simulate current out) | Current Value or simulated value (Simulate digital out) | Group mask | Data Sim | Switch Off Simulation Mode |  | Data simulation is on | In the diagnostis menu, switch off simulation mode
45 | 38 | DIAG_SIMU_ALARM | Coordinator | CDT_ALARM_CB_ALARM | 152 | CHECK_FUNCTION | CONFIG_STATUS |  |  | Group mask | Alarm Sim | Switch Off Alarm Simulation |  | Alarm simulation is on | In the diagnostis menu, switch off the alarm simulation mode | SRS,FEP300
46 | 39 | CURRENT_OUT_FIXED_MODE | Current Out | CURR_DIAG_FIXEDMODE | 154 | CHECK_FUNCTION | CONFIG_STATUS | fixed Value | No response | Group mask | Fixed Current Output | Check Simulation Mode or HART Address |  |  |  | SRS,FEP300
47 | 40 | CURRENT_OUT_FAILURE | Current Out | CURR_DIAG_SYSLIMIT_LO | 203 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Current Output Fault | Current out fault. Check Wring! |  |  |  | FEP300
48 | 40 |  | Current Out | CURR_DIAG_HWLIMIT_HI
49 | 40 |  | Current Out | CURR_DIAG_HWLIMIT_LO
50 | 40 |  | Current Out | CURR_DIAG_UNCALIB_FAILSAFE_LO
51 | 40 |  | Current Out | CURR_DIAG_INVALID_INPUT_VALUE
52 | 41 | CURRENT_READBACK_FAILED_HIGH | Electronic Services | ELETTRONICSERVICES_ALARM_CURRENT_READBACK_FAILED_HIGH | 217 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Output Read back High Failure | Current Output Calibrated? Circuit Broken? |  | Fault affecting transmitter electronics unit. | Replace the electronic unit or contact ABB Service. | FEP300
53 | 42 | CURRENT_READBACK_FAILED_LOW | Electronic Services | ELETTRONICSERVICES_ALARM_CURRENT_READBACK_FAILED_LOW | 216 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Output Read back Low Failure | Current Output Calibrated? Circuit Broken? |  | Fault affecting transmitter electronics unit. | Replace the electronic unit or contact ABB Service. | FEP300
54 | 43 | NV_REPLACE_WARN | Coordinator | CDT_ALARM_REPLACE_FE_TO_CB_BLOCKED | 54 | MAINTENANCE | OPERATING_CONDITION | Current Value | No response | Group mask | NV Replace Warning | NV data replace warning. Contact ABB Service
55 | 43 |  | Coordinator | CDT_ALARM_REPLACE_REQUIRED_FE_TO_CB_BLOCKED
56 | 43 |  | Coordinator | CDT_ALARM_REPLACE_REQUIRED_BOTH_DIRECTION_ENABLED
57 | 44 | FE_RAM_FAILURE | Coordinator | CDT_FE_RAM_FAILURE | 208 | FAILURE | HW_STATUS_ELECTRONICS | general alarm | Collective Alarm | no | Sensor RAM Fault | Replace the Front end Board |  | RAM damaged or RAM faults | Replace frontend board or contact ABB Service. | Merrick,FEP300
58 | 45 | TOT_COUNTER_STOP | Totalizer | TOTALIZER_STOP_ALARM | 155 | CHECK_FUNCTION | OPERATING_CONDITION | Current value | No response | Group mask | Totalizer Stop | Totalizer stop |  |  |  | DiagnosisInfo from Kevin, FEP300
59 | 47 | NO_HART_BURST_IN | Mapper | MAPPER_ALM_NO_HART_BURST_IN | 153 | CHECK_FUNCTION | CONFIG_STATUS | Current value | No response | Group mask | No HART Input | Check HART connection
Note on G6: Author: Not FAILURE because it needs to be masked, so selected as OFF_SPECIFICATION

## Alarm Output condition(only)
No. | subsystem | AlarmID. | Output Condition DESCRIPTOR | Priority | CLASSIFICATION | GROUP | remap | A | HMI Recovery Action | Remedy | Comments
1 | CDT | 0 | VORTEX_SENSOR_FAILURE | 213 | FAILURE | HW_STATUS_SENSOR | PV_SENSOR_FAILURE | Sig. Sensor Fault | Contact ABB Service |  Check the connection between the piezo sensor and front end board and establish the connection if necessary  Check that the three sensor leads are insulated from each other and from ground (dc resistance > 50 MegOhm). The capacitance between the yellow or red leads to the blue lead should be between 3.5 – 4.5 nF (for high-temperature sensors approx. 500 pF). If resistance or capacitance are not in the range, contact ABB service to replace the sensor.  If above two solutions are failed, replace the front end board or contact ABB service.
2 | CDT | 1 | TEMP_SENSOR_FAILURE | 212 | FAILURE | HW_STATUS_SENSOR | SV_SENSOR_FAILURE | Int. T Sesnor Fault | Contact ABB Service | Check the connection between the PT100/PT1000 sensor and front end board and establish the connection if necessary | need to be masked,so classification to be off-spec,keep priority and group classification
3 | CDT | 2 | VIBRAT_SENSOR_FAILURE | 211 | FAILURE | HW_STATUS_SENSOR | TV_SENSOR_FAILURE | Vbr. Sesnor Fault | Contact ABB Service |  Check the connection between the piezo sensor and front end board and establish the connection if necessary  Check that the three sensor leads are insulated from each other and from ground (dc resistance > 50 MegOhm). The capacitance between the yellow or red leads to the blue lead should be between 3.5 – 4.5 nF (for high-temperature sensors approx. 500 pF). If resistance or capacitance are not in the range, contact ABB service to replace the sensor.  If above two solutions are failed, replace the front end board or contact ABB service.
4 | AnalogInput | 3 | ANALOGINPUT_OUT_OF_RANGE | 106 | OFF_SPECIFICATION | OPERATING_CONDITION | FV_SENSOR_FAILURE | AI Out of Range | Check AI Current
5 | MVM/SVM | 4 | FLOWRATE_REACH_QMAX | 114 | OFF_SPECIFICATION | OPERATING_CONDITION | PV_SENSOR_OUT_LIMS | Max Alarm Flowrate | Check Flowrate. Check Limits Setting | Reduce the flowrate or increase the value for the max alarm.
6 | CDT | 5 | PROCESS_TEMP_REACH_TMAX | 112 | OFF_SPECIFICATION | OPERATING_CONDITION | SV_SENSOR_OUT_LIMS | Max Alarm Int. Temp | Check Temperature. Check Limits Setting | Check process, adjust alarm limits if required.
7 | AnalogInput | 6 | ANALOGINPUT_CUT_OFF | 107 | OFF_SPECIFICATION | OPERATING_CONDITION | TV_SENSOR_OUT_LIMS | External Output Cutoff | Check AI Current. Check AI Setting
8 | Mapper | 7 | PRESSURE_REACH_PMAX | 102 | OFF_SPECIFICATION | OPERATING_CONDITION | FV_SENSOR_OUT_LIMS | Max Alarm Pressure | Check Pressure. Check Limits Setting | Check process, adjust alarm limits if required.
9 | MVM/SVM | 10 | FLOWRATE_REACH_QMIN | 113 | OFF_SPECIFICATION | OPERATING_CONDITION | PV_SENSOR_OUT_RANGE | Min Alarm Flowrate | Check Flowrate. Check Limits Setting | Increase the flowrate or increase the value for the max alarm.
10 | CDT | 11 | PROCESS_TEMP_REACH_TMIN | 111 | OFF_SPECIFICATION | OPERATING_CONDITION | SV_SENSOR_OUT_RANGE | Min Alarm Int. Temp | Check Temperature. Check Limits Setting | Check process, adjust alarm limits if required.
11 | Current out | 12 | CURRENT_OUT_SATURATED | 108 | OFF_SPECIFICATION | OPERATING_CONDITION | TV_SENSOR_OUT_RANGE | Current Output Saturated | Check the Primary Variable. Check the Saturation Limit | Adjust the Saturation Limit or the working range if possible.
12 | Mapper | 13 | PRESSURE_REACH_PMIN | 101 | OFF_SPECIFICATION | OPERATING_CONDITION | FV_SENSOR_OUT_RANGE | Min Alarm Pressure | Check Pressure. Check Limits Setting | Check process, adjust alarm limits if required.
13 | CDT | 16 | SNR_FAILURE | 210 | FAILURE | HW_STATUS_ELECTRONICS | FE_BOARD_FAILURE | Bad SNR | Increase the Flow. Check Limits Setting | Increase the flowrate or adjust alarm limits if required.
14 | CDT | 17 | FE_NV_FAILURE | 209 | FAILURE | HW_STATUS_ELECTRONICS |  | Sensor NV Error | Replace the Front end Board | Replace the front end board or contact ABB service.
15 | CDT | 18 | SENSOR_CAL_STATUS | 51 | MAINTENANCE | OPERATING_CONDITION | FE_NV_WARNING | Sensor Not Calibrated | Contact ABB Service | Sensor is uncalibrated or calibration status is not set to "calibrated".
16 | CDT | 19 | FE_SYNC_MISSING | 214 | FAILURE | HW_STATUS_ELECTRONICS | FE_BOARD_NOT_DETECTED | Sync. Signal Error | Contact ABB Service | Power off and on the transmitter and check if the error persists. If yes replace the front end board as soon as possible.
17 | FrontendInterface | 20 | FE_BOARD_COMM_ERROR | 215 | FAILURE | HW_STATUS_ELECTRONICS |  | Sensor Comm Error | Check Wiring | Check the connection between front end board and communication board or contact ABB service.
18 | CDT | 23 | CB_NV_FAILURE | 207 | FAILURE | HW_STATUS_ELECTRONICS |  | Transmitter NV Error | Replace the Communication Board | Replace the communication board or contact ABB service.
19 | AI | 24 | AI_SPI_COM_FAIL | 202 | CHECK_FUNCTION | CONFIG_STATUS | CB_NV_WARNING | AI Comm Error | Check Wiring | Check the connection between IO board and communication board or contact ABB service.
20 | DO | 25 | DIGITAL_OUT_CUTOFF | 103 | OFF_SPECIFICATION | OPERATING_CONDITION | NV_STORAGE_ACTIVE | Pulse Output Cutoff | Check Pulse Output Config | In the "Easy Set-up" menu, reduce the "Pulses per unit" value.
21 | SVM | 26 | RE_OUT_LINEAR_RANGE | 109 | OFF_SPECIFICATION | OPERATING_CONDITION | NV_CONCISTENCY_CHECK | Re. out of Range | Increase The Flow | Increase the flow.
22 | MVM | 30 | STEAM_TYPE_MISMATCH | 116 | OFF_SPECIFICATION | OPERATING_CONDITION |  | Wrong Steam Type | Check Steam Type Setting
23 | CDT | 31 | MAINTENANCE_WARNING | 52 | MAINTENANCE | OPERATING_CONDITION |  | Maintenance Warning | Contact ABB Service | Adjust the maintenance cycle if required or contact ABB service to recalibrate the device
24 | Electronic Services | 32 | CB_INSUF_INPUT_VOLTAGE | 53 | MAINTENANCE | OPERATING_CONDITION |  | Voltage Warning | Check Power Supply | Check the voltage at the terminal block, and if it is not within the valid range, check the external power supply
25 | Electronic Services | 33 | AMBIENT_TEMP_REACH_TMIN | 104 | OFF_SPECIFICATION | OPERATING_CONDITION |  | Min Alarm Housing Temp | Check Housing Temp. Check Limits Setting | Check whether the device has been correctly insulated Increase the ambient temperature.
26 | Electronic Services | 34 | AMBIENT_TEMP_REACH_TMAX | 105 | OFF_SPECIFICATION | OPERATING_CONDITION |  | Max Alarm Housing Temp | Check Housing Temp. Check Limits Setting | Check whether the device has been correctly insulated Reduce the ambient temperature.
27 | SVM | 35 | LOW_FLOWRATE_CUTOFF | 110 | OFF_SPECIFICATION | OPERATING_CONDITION |  | Flowrate Cutoff | Check Flowrate. Check Low Flow Cutoff Setting | Increase the flow or adjust low flow cut off setting in the “Easy Set-up -> Low Flow Cut off” menu
28 | MVM/SVM | 36 | FLOWRATE_BG_103 | 115 | OFF_SPECIFICATION | OPERATING_CONDITION |  | Flowrate > 103% | Check Flowrate. Check Limits Setting | Increase the flow range end value in the "Easy Set-up -> Qmax" menu.
29 | CDT | 37 | DATA_SIMU_ALARM | 151 | CHECK_FUNCTION | CONFIG_STATUS |  | Data Sim | Switch Off Simulation Mode | In the "Diagnostics" menu, deactivate simulation mode.
30 | CDT | 38 | DIAG_SIMU_ALARM | 152 | CHECK_FUNCTION | CONFIG_STATUS |  | Alarm Sim | Switch Off Alarm Simulation | In the "Diagnostics" menu, deactivate alarm simulation
31 | Current out | 39 | CURRENT_OUT_FIXED_MODE | 154 | CHECK_FUNCTION | CONFIG_STATUS |  | Fixed Current Output | Check Simulation Mode or HART Address | Deactivate simulation mode in the "Process Alarm" menu or set the HART address to 0 in the "Communication" menu.
32 | Current out/Electronic | 40 | CURRENT_OUT_FAILURE | 203 | FAILURE | HW_STATUS_ELECTRONICS |  | Current Output Fault | Calibrated? Well Configured? | Perform an Output Trimming and if the error persists the communication board must be replaced. Check the device configuration.
33 | Electronic Services | 41 | CURRENT_READBACK_FAILED_HIGH | 217 | FAILURE | HW_STATUS_ELECTRONICS |  | Output Read back High Failure | Current Output Calibrated? Circuit Broken? | A DAC (digital to output converter) trimming should be performed and if the error persists the communication board must be replaced | can not be displayed because of device reset
34 | Electronic Services | 42 | CURRENT_READBACK_FAILED_LOW | 216 | FAILURE | HW_STATUS_ELECTRONICS |  | Output Read back Low Failure | Current Output Calibrated? Circuit Broken? | A DAC (digital to output converter) trimming should be performed and if the error persists the communication board must be replaced | can not be displayed because of device reset
35 | CDT | 43 | NV_REPLACE_WARN | 54 | MAINTENANCE | OPERATING_CONDITION |  | NV Replace Warning | Check Replace Config | The replacement operation must be executed: 1. Move the SW 1.2 at enable replace mode ; 2. Select the SW 1.3 at the right direction. 3. After HMI is powered on, move SW1.2 at disable replace mode.
36 | CDT | 44 | FE_RAM_FAILURE | 208 | FAILURE | HW_STATUS_ELECTRONICS |  | Sensor RAM Fault | Replace the Front end Board | Replace the front end board or contact ABB service.
37 | TOT | 45 | TOT_COUNTER_STOP | 155 | CHECK_FUNCTION | OPERATING_CONDITION |  | Totalizer Stop | Totalizer stop | Start totalizer in “Totalizer->Operation->Start” menu.
39 | Mapper | 47 | NO_HART_BURST_IN | 153 | CHECK_FUNCTION | CONFIG_STATUS |  | No HART Input | Check HART connection